## Slide 1: RENASYS™ GO NPWT Clinical training
Notes: The RENASYS GO device is designed to provide NPWT to a closed environment over a wound, in order to evacuate exudates from the wound site to
a disposable container, which may promote wound healing via removal of fluids, including irrigation and body fluids, wound exudates and infectious materials.
Todays presentation will cover the following agenda (Next slide)

## Slide 2: Objectives
Indications, contraindications, warnings, precautions
Product overview
Functionality
RENASYS™ GO alarms
Dressing application – foam and gauze
Bridging
Changing canister
Dressing removal
Support

## Slide 3: Indications for use:
RENASYS™ GO is indicated for patients who would benefit from a suction device (Negative Pressure Wound Therapy), as it may promote wound healing via removal of fluids, including irrigation and body fluids, wound exudates and infectious materials
Appropriate wound types include:
Chronic
Acute
Traumatic
Sub-Acute and dehisced wounds
Ulcers (such as pressure or diabetic)
Partial-thickness burns
Flaps and grafts
3
Notes: Speaker notes/guidance:

## Slide 4: Contraindications
The use of the RENASYS™ GO System is contraindicated in the presence of:
Necrotic tissue with eschar
Untreated osteomyelitis
Malignancy in wound (with exception of palliative care to enhance quality of life)
Exposed arteries, veins, organs or nerves
Non-enteric and unexplored fistulas
Anastomotic sites
4
Notes: Speaker notes/guidance:

## Slide 5: Warnings
1. Carefully monitor patients for signs of bleeding, which may lead to interruption in therapy and hemodynamic instability. If such symptoms are observed, immediately discontinue therapy, take appropriate measures to control bleeding, and contact treating clinician
2. Patients suffering from difficult hemostasis or who are receiving anticoagulant therapy have an increased risk of bleeding. During therapy, avoid using hemostatic products that, if disrupted, may increase the risk of bleeding
3. Do not use directly on exposed blood vessels or organs. Sharp edges such as bone fragments must be covered or removed prior to initiating therapy, due to risk of puncturing organs or blood vessels drawn closer under the action of negative pressure
4. NPWT has not been studied on pediatric patients. Patient size and weight should be considered when prescribing the device
5. Foam or gauze must not be tightly packed or forced into any wound area. Over-packing may interfere with distribution of NPWT evenly across the wound. This may decrease the ability of the wound to properly contract and permit exudate to remain in wound
6 In the event defibrillation is required, disconnect device from wound dressing prior to defibrillation. Remove wound dressing only if its location will interfere with defibrillation
7. Device is not MRI compatible. Do not bring device into MRI suite. Prior to entering MRI suite, disconnect device from dressing. Dressing can remain intact on patient
8. Do not use the device in oxygen enriched environment or in areas where there is danger of explosion (e.g., hyperbaric oxygen unit).
9. When operating, transporting or disposing of device and accessories, there is risk of infectious liquids being aspirated or contamination of device assembly through incorrect use. Universal precautions should be observed whenever working with potentially contaminated components or equipment
10. Device and canister kits are provided non-sterile and should not be placed within a sterile field
5
NOTE: Full device operation is found in the User Manual for the RENASYS™ GO Device
Notes: Speaker notes/guidance:

## Slide 6: Precautions
  Suffering from infected blood vessels
  Receiving anticoagulant therapy or platelet aggregation inhibitors, in addition to patients with intrinsic coagulation problems such as low platelet counts
  Actively bleeding or have friable blood vessels or organs
  Suffering from abnormal wound hemostasis
  Untreated for malnutrition
  Noncompliant or combative
  Suffering from wounds in close proximity to blood vessels or friable fascia
6
NOTE: Full device operation is found in the User Manual for the RENASYS™ GO Device
When monitoring patients for delivery of therapy, ensure wound dressing is free of air leaks, fully compressed and firm to the touch
1. More frequent device and wound dressing monitoring, should be taken for patients who are or may be:
Notes: Speaker notes/guidance:

## Slide 7: Precautions continued
2. As a condition of use, device should only be used by qualified and authorized personnel. User must have necessary knowledge of the specific medical application for which NPWT is being used
3. For patients with high risk of bleeding, use 300mL canister. Ensure the 300mL canister viewing is checked frequently for signs of bleeding
4. The dressing seal may be lost or pooling may occur, without alarm activation, when an occlusion forms on the wound side of the dressing. Viscous, purulent or serosanguineous drainage may contribute to occlusion of the dressing. Regular monitoring of device and dressing is required to ensure full delivery of therapy and exudate removal. Ensure the wound dressing is free of air leaks, fully compressed and firm to the touch whenever therapy is active
5. Underlying structures, such as bone, tendons, ligaments and nerves should be covered with natural tissue or a non-adherent dressing layer prior to applying the NPWT dressing to ensure protection and minimize the risk of damage from direct contact with the dressing
6. To minimize the risk of bradycardia, do not place NPWT in proximity to the vagus nerve
7. In the event a patient with spinal cord injury experiences autonomic dysreflexia, discontinue use of NPWT and immediately seek medical assistance
8. When treating enteric fistulas, do not place NPWT dressing in direct contact with exposed bowel. Cover the wound bed, including fistula opening, with non-adherent gauze or with one layer of saline moistened gauze. During the course of treatment patient’s fluid levels must be closely monitored
7
NOTE: Full device operation is found in the User Manual for the RENASYS™ GO Device
Notes: Speaker notes/guidance:

## Slide 8: Precautions continued
9. Avoid use of circumferential dressings except in cases of edema or heavily exuding extremities, where this technique may be necessary to maintain a seal. Consider using multiple drapes to minimize risk of decreased distal circulation. Regularly assess distal pulses, and discontinue therapy if changes in circulation are detected
10. Monitor patient for any signs of local or systemic infection. Infected wounds may require more frequent dressing changes. If there are any signs of systemic infection or advancing infection at wound area, contact treating clinician immediately
11. If multiple pieces of foam or gauze are needed to fill the wound profile, count and record how many pieces are present to ensure all pieces are removed at a dressing change to minimize the risk of retention and possible infection
12. NPWT should remain on for duration of treatment. The length of time a patient may be disconnected from device is a clinical decision based on individual characteristics of patient and wound. Factors to consider include location of wound, volume of drainage, integrity of dressing seal, assessment of bacterial burden and patient’s risk of infection
13. Do not use a dressing kit with breached or damaged packaging
14. Use of NPWT presents a risk of tissue ingrowth. Tissue ingrowth may be reduced by reducing therapy pressure, using a wound contact layer or increasing the frequency of dressing changes
15. NPWT should not be painful. If patient reports discomfort, consider reducing pressure setting and use of a wound contact layer. Pressure setting is a clinical decision based on individual characteristics of patient and wound. Factors to consider include location of wound, volume of drainage and integrity of dressing seal
8
NOTE: Full device operation is found in the User Manual for the RENASYS™ GO Device
Notes: Speaker notes/guidance:

## Slide 9: Precautions continued
16. Maintain regular monitoring of device and wound site during therapy to ensure therapeutic treatment and patient comfort
17. Device is only to be used with Smith & Nephew authorized components. Use of any other products have not been proven safe and effective with RENASYS™ GO device
18. Ensure canister tubing and RENASYS Soft Port are installed completely and without any kinks to avoid leaks or blockages in vacuum circuit. Position device and tubing appropriately to avoid risk of a trip hazard. Device and system tubing should be positioned in a more prominent location, no more than 19in or 50cm higher than the wound to ensure optimization of therapy and prevent therapy interruption
19. When bathing or showering patient must disconnect from device, protecting both ends of RENASYS Soft Port tubing using tethered caps. Ensure aeration disc located near Quick Click Connector is free of moisture before reactivation of therapy to ensure proper alarm functionality and prevent interruption in therapy
20. If any liquids penetrate device, discontinue use and return to your Smith & Nephew authorized provider for service
21. CT scans and x-ray have the potential to interfere with some electronic medical devices. Keep device out of x-ray or scanner range
22. Do not use in the presence of a flammable anesthetic mixture with air or with oxygen or nitrous oxide
23. Electrical power can only be removed by disconnecting power cord or AC power adaptor. Take care in positioning the device to allow access to cord receptacle
9
NOTE: Full device operation is found in the User Manual for the RENASYS GO Device
Notes: Speaker notes/guidance:

## Slide 10: Precautions continued
24. If power cord is damaged, wires are frayed or exposed, do not use power cord. Contact your Smith & Nephew representative for a replacement cord
25. Canisters should be changed at least once a week, whenever there is a change in patient or in the event that canister contents reach maximum volume indication (300mL or 750mL fill line)
Do not wait for the canister over-capacity alarm to activate to change canister
26. Canisters kits are single use devices. Do not reuse
27. Do not apply SECURA™ NO-STING SKIN-PREP barrier film wipes directly to open wounds. SECURA NO-STING SKIN-PREP barrier film is flammable. Use in a well ventilated area. Avoid using around flames and sources of ignition Keep out of reach of children. For external use only
28. As with all adhesive products, apply and remove dressing carefully from sensitive or fragile skin to avoid blistering and skin stripping, especially after frequent dressing changes. Use of skin sealant may assist with protection of periwound skin
29. If patient must be disconnected, the ends of the RENASYS™ Soft Port and canister tubing should be protected using tethered caps to avoid leakage of fluid and cross contamination
10
NOTE: Full device operation is found in the User Manual for the RENASYS GO Device
Notes: Speaker notes/guidance:

## Slide 11: Physician orders
Prior to placement of RENASYS™ GO, the medical professional treating the wound must assess how to best use the system for an individual wound. It is important to carefully assess wound and patient to ensure clinical indications for Negative Pressure Wound Therapy (NPWT) are met
All orders should include:
• Wound location, size and type
• Smith & Nephew wound dressing kit
• Pressure settings
• Frequency of dressing changes
• Adjunctive dressings
11
NOTE: Full device operation is found in the User Manual for the RENASYS GO Device
Notes: Speaker notes/guidance:

## Slide 12: RENASYS™ GO Product overview
Notes: The RENASYS GO device is designed to provide NPWT to a closed environment over a wound, in order to evacuate exudates from the wound site to
a disposable container, which may promote wound healing via removal of fluids, including irrigation and body fluids, wound exudates and infectious materials.
Todays presentation will cover the following agenda (Next slide)

## Slide 13: RENASYS™ GO overview
Vacuum: -40 to -200 mmHg
The device adjustments can be made in 10mmHg increments between 40mmHg and 100mmHg
20mmHg increments between 100mmHg and 200mmHg
Dimensions: 175mm x 210mm x 85mm
Weight: 2.4 lb (1.1kg)
Battery operation: Up to 20 hours
Battery operations: Lithium ion rechargeable
Charging time: Approx. 3 hours for 80% capacity
Canisters: 300mL and 750mL

## Slide 14: RENASYS™ GO user interface
Power button
Front view
Start/pause therapy and select
Keypad lock
Up and down selectors
Battery status indicator
Display screen
Status/alarm indicator
Audio pause
Canister
Fill warning label

## Slide 15: Status lights
Status LED:
UNLIT = Standby
Green = Therapy
= Alarm condition
Yellow
Notes: The status LED of the device has three levels of display
Unlit – the device is on standby
Green – the device is in therapy mode
Yellow – the device has an alarm condition requiring investigation

## Slide 16: Starting therapy
Select the prescribed pressure setting using Up and Down buttons
Press Select button to start therapy
Finished dressing should be fully compressed, firm to the touch and leak free
Setting pressure level is a decision that the healthcare provider must make based on an individual assessment of the particular wound
40-120mmHg is the recommended therapeutic pressure range
In the event the wound is difficult to seal, the pressure setting may be increased to achieve the desired negative pressure at the wound bed
Device will show set pressure level. If device detects operation outside of the set pressure, it will result in alarm activation
Caution: Before starting therapy, ensure the device and system tubing are located no more than 19 inches (50cm) or higher than the wound and is away from any direct sources of heat

## Slide 17: Lock or unlock keypad
To lock the user interface when device is active, press and hold Keypad Lock button for 2 seconds
When activated, the light will illuminate solid blue
If the keypad on user interface has not been used for 15 minutes, key pad will automatically lock
In active mode, if keypad or user interface has not been used for 15 minutes, keypad will automatically lock
To unlock the user interface, press and hold the Keypad Lock button for 2 seconds
Note: Locking keypad only locks the row of keys under display screen. The Power Button is not locked

## Slide 18
Inlet connector and “O” ring
Canister clips
Shoulder Strap
Side view
Electrical power inlet
Power Supply
Notes: The other key features of the GO pump include a Shoulder strap. The shoulder strap is used for carrying the device by the patient for mobility. The strap is also used in the Acute care setting to hang the pump on a bed rail or IV pole to ensure its vertical orientation. The inlet connector and o-ring which creates a seal with the canister to ensure negative pressure in the system; Canister clips that securely lock the canister to the pump housing; electrical power inlet and power supply.

## Slide 19: RENASYS™ GO orientation
Recommended orientation
Incorrect orientation
Notes: RENASYS GO device may also be used in a horizontal orientation, with device facing upward. However, to maximize canister volume and optimize canister over-capacity alarm, upright device orientation is recommended.

## Slide 20: RENASYS™ GO orientation
RENASYS GO may also be used in a horizontal orientation, with the device facing upward. However, to maximize canister volume and optimize canister over-capacity alarm, upright device orientation is recommended
Caution: When device is operating in horizontal orientation, facing upward, alarm functionality may be compromised. Monitor wound dressing to ensure it is fully compressed and firm to the touch
Notes: Caution: When device is operating in horizontal orientation, facing upward, alarm functionality may be compromised. Monitor wound dressing to ensure it is fully compressed and firm to the touch

## Slide 21: Device/tubing height relative to the wound
RENASYS™ GO should not be placed greater than 19in (50cm) above the wound
< 19in (50cm)
Ensures that the appropriate vacuum (pressure) is transmitted to the wound site
Notes: NOTE: Positioning the RENASYS GO at the wound level and reasonable distances below the wound level is fine. Ensuring the device is not positioned greater than 19 in (50cm) above the wound is because the weight of fluid in a long vertical run of vertical tube may reduce the effective vacuum at the wound.

## Slide 22: RENASYS™ GO canister selection and installation
RENASYS GO is designed to be used with the RENASYS GO 300mL or 750mL canisters
300mL Canister
750mL Large canister
For patients with high risk of bleeding, use 300mL canister. Ensure the 300mL canister viewing is checked frequently for signs of bleeding
Notes: For patients with high risk of bleeding, use 300ml canister. Ensure the 300ml canister viewing is checked frequently for signs of bleeding

## Slide 23: Modes of operation
To change therapy between continuous and intermittent mode:
Turn device off
Simultaneously press and hold Down + Select + Power buttons for 2 seconds
Use Up or Down buttons to scroll between continuous and intermittent mode. Press the Select button to confirm
Once therapy is started, display screen will show which mode of therapy has been selected

## Slide 24: RENASYS™ GO device - alarms
Over-vacuum
High vacuum
Leak alarm
Low vacuum
Complete blockage/canister over-capacity alarm
Inactive alarm
Low battery alarm

## Slide 25
RENASYS™ Dressing Kits
RENASYS-F Foam Dressing Kit with Soft Port
  Recognized industry standard; Proven efficacy
RENASYS-G Gauze Dressing Kit with Soft Port
  Enhanced patient comfort
  High conformability
Bringing choice and flexibility to NPWT by helping to promote wound healing and improve patient comfort, while allowing clinicians to use their full clinical judgment for each individual patient
Notes: The extensive line of Smith & Nephew RENASYS* NPWT Dressing kits is bringing choice and flexibility to Negative Pressure Wound Therapy by helping to promote wound healing and improve patient comfort, while allowing clinicians to use their full clinical judgment in treating each individual patient.
Foam
Recognized industry standard.
Foam is the most commonly used dressing in Negative Pressure Wound Therapy. It is easy to apply and appropriate for a wide range of wound types and sizes.
Proven efficacy.
Foam is proven to address the common treatment goals associated with negative pressure, including reductions in wound dimensions, exudate levels and wound odor, and an improvement in the quality of the wound bed.1
Gauze
Enhanced patient comfort -Gauze is ranked number one in patient comfort upon application and removal, and is suitable for most wound types and sizes.
High conformability -Gauze is moldable and forgiving, making it easy to apply to any wound. It is especially suited to anatomically challenging locations with abnormal contours or undermining, such as the sacrum. It is also ideal for large surface areas, explored fistulas and tunneling wounds.

## Slide 26: RENASYS™ Soft Port
Clinically proven and powerfully enhanced
Reduces the need for bridging1, 2
Enables management of a wide range of wound types and exudate levels with a single application technique 1
Device does not block throughout the duration of therapy, despite compressions, twists or folds in the line 1
Comfortably effective
May reduce the risk of pressure-related injuries and eases patient discomfort by replacing hard tubing with a soft, cushioned channel1
1 Hudson D. et. al. Evaluation od a new Negative Pressure Wound Therapy (NPWT) suction Port (RENASYS Soft Port). Poster presented EWMA, Copenhagen 2013.
2 Schwartz J, et al. An open, prospective, clinical evaluation to determine the clinical efficacy of a new Negative Pressure Wound Therapy (NPWT) system with soft port technology after partial diabetic foot amputation. Poster presentation, 2013, SAWC.

## Slide 27: Choosing an NPWT wound filler
Factors to be considered include
    Wound size and volume
    Contour of wound bed
    Amount and type of exudate
    Patient comfort and preference
    Characteristics of granulation tissue
    Caregiver skills

## Slide 28: NPWT wound filler and wound interface
Definitions:
Wound fillers – primarily gauze and foam function to:
  Fill the wound cavity
  Allow the negative pressure to be transmitted to the wound surface
  Function as a conduit for wound fluid to be transferred to suction source
Wound interface – Primarily non-adherent layer or ACTICOAT™ Flex
  Used to protect fragile structures from direct contact with negative pressure
  Use to minimize tissue ingrowth with foam filler
  ACTICOAT Flex may be used if there is actual or potential increased bacterial
  burden

## Slide 29: Choosing an NPWT wound filler
Assessing wound bed characteristics:
  Is a wound contact layer needed?
  Assess the wound for infection. Is a topical antimicrobial needed?
  If wound is infected, it could require more frequent dressing
  changes and the patient should be monitored closely
Selecting a wound filler/interface:
  Wound bed contour. Are there tunnels and undermining?
  Assessment of wound exudate (levels and characteristics)
  Patient comfort
  Assessment of granulation tissue
  Caregiver skills
Note: The use of negative pressure presents a risk of tissue ingrowth into the foam/gauze. Tissue ingrowth may be reduced by lowering therapy pressure, using a wound contact layer or by increasing the frequency of dressing changes

## Slide 30: Choosing a NPWT wound filler – wound bed contour
Wound bed with irregular contour.
Foam wound filler may not intimately
contact irregular shape spaces in
wound bed.
Gauze wound filler easily maintains
contact with irregular surface.
Do not tightly pack or force foam into any areas of the wound. Do not place foam into blind or unexplored tunnels. The use of negative pressure presents a risk of tissue ingrowth into the foam. Tissue ingrowth may be reduced by reducing therapy pressure, using a wound contact layer or by increasing the frequency of dressing changes
Notes: Gauze wound fi ller easily maintains contact with an irregular surface. Foam wound fi ller may
not intimately contact irregular shape spaces in wound bed. In this case, gauze and foam
may be used in combination

## Slide 31: Choosing a NPWT wound filler – wound exudate
The choice of wound filler will be influenced by the amount and consistency of wound exudate
Small
Large
Wound
Exudate
Level
Gauze
Foam
Note:
In wounds with large amounts of exudate a wound interface (non-adherent layer) is
generally not recommended
If using gauze with larger amounts of exudate the drain should be placed close to the
wound bed over a single layer of gauze
Notes: Precautions specific to gauze
Use of gauze as a filler in wounds that are in a weight bearing location or have moderate to
heavy drainage may challenge the transfer of fluid and vacuum. Foam is recommended in
these wounds.
RENASYS◊ Foam and Gauze Fillers may be combined within the same wound when tunneling
or undermining is present. Gauze may be used in the areas of undermining or tunneling,
with foam placed in the remainder of the wound cavity. In wounds with large amounts of
exudate a wound interface (non-adherent layer) is generally not recommended.
Ensure gauze is placed into the contours of the wound rather than tightly packed into the
wound. This will aid fluid transference.

## Slide 32: Choosing a NPWT wound filler – patient comfort
The choice of wound filler will be influenced by the amount of pain the patient experiences during NPWT and with dressing changes
Pain is a very subjective experience and will vary with each patient. Research has validated that patients report less pain with gauze1
Lower level
Higher level
Pain/discomfort
Gauze
Foam
1 Grothier. et. al. The level of pain experienced during therapy and dressing changes of gauze-based Negative Pressure Wound therapy. Presented at Wounds UK 2009

## Slide 33: RENASYS™-F Foam Dressing Kit with Soft Port application procedure
Contents of foam kit:
Polyurethane Foam
Soft Port
Transparent Films

## Slide 34: RENASYS™-F Foam Dressing Kit with Soft Port
Cut the foam dressing to fit the size and shape of the wound.
Do not cut dressing directly over wound as fragments may fall
into wound
Place foam into wound cavity
Foam should completely fill the wound cavity
It may be necessary to stack pieces of foam in deep wounds
Caution: Do not over pack foam into the wound
A non-adherent dressing may be applied to the wound prior
to placing the foam into the wound bed if required

## Slide 35: Contact layers
Cut Non-Adherent Gauze to fit wound bed
Place a single layer across wound bed
ACTICOAT™ Flex and ACTICOAT may be used as alternate contact layers for up to 3 days
ACTICOAT must be fenestrated
Notes: ACTICOAT and ACTICOAT Flex come in 3 and 7 day products. Both products may be used in wounds with high risk for infection or wounds that are already infected.

## Slide 36: RENASYS™-F Foam Dressing Kit with Soft Port
Cut the foam dressing to fit the size and shape of the wound and place the cut foam into the wound
Avoid over packing. Foam should
completely fill the wound cavity.
It may be necessary to stack
pieces of foam in deep wounds
Notes: Note:
Prepare the wound bed by cleansing per hospital protocol. Pat dry, and apply skin prep to periwound area.
Be sure not to over pack the wound, this could cause damage to the tissue in the wound bed.
Always cut the foam away from the wound bed (you do not want pieces of foam to fall into the wound).
It is important to apply enough foam to the wound bed to offset the amount of contraction of the foam after suction has been applied.
The foam will have to be above the level of the wound in order to be at skin level after suction is applied.
It is recommended that the number of pieces of foam be recorded in the wound documentation or on the dressing film to ensure all foam is removed at dressing change.
Always use the minimum number of pieces of foam possible.

## Slide 37: RENASYS™-F Foam Dressing Kit with Soft Port
Cover foam with Transparent Film
Film should extend at least 2in (5.08cm) beyond wound margin to facilitate adequate seal

## Slide 38: RENASYS™-F Foam Dressing Kit with Soft Port
Cut a circular opening in film no less than 2.0cm (0.79in) in diameter
Remove any excess trimmed film
Notes: The Soft Port opening is actually 0.6 in (1.5cm) diameter.

## Slide 39: Remove the RENASYS™ Soft Port liner
Remove the liner from the Soft Port dressing
Align Soft Port hole directly over film hole
Notes: Note:
Do not cut the Transparent Film surrounding the Port. It will be very difficult to apply the Port without the top carrier lining/frame to hold/guide the dressing. The top carrier lining/frame gives the dressing structure/stability for easy application.

## Slide 40: Apply and anchor the RENASYS™ Soft Port
Use gentle pressure to anchor the Soft Port dressing to the Transparent Film
Smooth the dressing down while removing the frame
Notes: There is a pre-cut slit under the Soft Port “tail” to help you easily remove the stabilization frame.

## Slide 41: Quick Click Connector
Connect the RENASYS™ Soft Port to canister using Quick Click Connector
Do not tape over or otherwise occlude aeration disc
Notes: Quick Click snaps together, want to hear an audible “click”
Put caps together, keeps them clean and out of the way
Taping over aeration disc can give a false blockage alarm. It allows air into the system, want to keep it free of occlusions

## Slide 42: Finished foam dressing
Finished and leak-free dressing should be firm to the touch
Routine inspection of the NPWT device and dressing should include visual inspection and palpation of the dressing
Carefully monitor patients for signs of bleeding, which may lead to interruption in therapy and hemodynamic instability. If such symptoms are observed, immediately discontinue therapy, take appropriate measures to control bleeding, and contact treating clinician
Notes: NOTE:
Foam may only be in wound bed for 2 hours without NPWT in use. If unable to troubleshoot pump, please remove foam within 2 hours and place a moistened normal saline gauze dressing. Notify prescribing clinician that dressing was removed. Do NOT just turn pump off and leave foam dressing in place.
Always assess wound with every assessment and turn. Look for firm “raisin like” appearance.
Pressure settings are patient and wound specific, to be determined by clinician.

## Slide 43: RENASYS™-G Gauze Dressing Kit with Soft Port application procedure
Contents of gauze kit:
Antimicrobial gauze
Non-adherent layer
Saline bullet
NO-STING SKIN-PREP™
Transparent Film
Ruler
Soft Port
Notes: Antimicrobial gauze (PHMB, Polyhexamethylene Biguanide). Does not kill bacteria, resists colonization (bacteriostatic not bactericidal like ACTICOAT

## Slide 44: RENASYS™-G Gauze Dressing Kit with Soft Port
Fill wound bed with saline moistened antimicrobial gauze
Cover with Transparent Film. Film should extend at least 2in (5.08cm) beyond wound margin to facilitate adequate seal
Cut a circular opening in film no less than 2.0cm in diameter
Remove any excess trimmed film.
Align Soft Port hole directly over film hole
Notes: Again, do not overfill or pack dressing into wound. Can cause more damage to wound bed.
Neither foam or gauze should ever touch the good skin that is surrounding the wound bed.

## Slide 45: Apply and anchor the RENASYS™ Soft Port
Use gentle pressure to anchor the Soft Port dressing to the Transparent Film
Smooth the dressing down while removing the frame

## Slide 46: Finished gauze dressing
The finished dressing should be leak-free and firm to the touch
Routine inspection of the NPWT device and dressing should include visual inspection and palpation of the dressing
Carefully monitor patients for signs of bleeding, which may lead to interruption in therapy and hemodynamic instability. If such symptoms are observed, immediately discontinue therapy, take appropriate measures to control bleeding, and contact treating clinician
Notes: Pressure settings are patient and wound specific, to be determined by clinician.

## Slide 47: Bridging techniques
Bridging technique – this technique is used to join two wounds that are close in proximity and/or to position the RENASYS™ Soft Port in an area away from the wound
Bridging can be done using both the RENASYS-G and RENASYS–F dressing kits
Consider bridging two wounds together using a wound filler for wounds separated by a distance less than 25cm. If the wounds are greater than 25cm consider using two RENASYS pumps

## Slide 48: Bridging – step 1
Protect the intact skin between the wounds with Transparent Film
Cut foam to the size and shape of wounds to be bridged
Begin securing with Transparent Film to hold foam in place

## Slide 49: Bridging – step 2
Cut additional foam and place on top of Transparent Film to form the bridge
Important: Contact must be made between the bridge foam and the foam in the wound bed (overlap)
NOTE: To optimize transfer of fluid and vacuum through the bridge, the foam dressing is recommended. In the instance where gauze is used, particularly for large volume and surface area wounds with moderate to heavy drainage, more frequent inspection of the dressing and dressing changes may be required
Notes: NOTE: To optimize transfer of fluid and vacuum through the bridge, the foam dressing is recommended. In the instance where gauze is used, particularly for large volume and surface area wounds with moderate to heavy drainage, more frequent inspection of the dressing and dressing changes may be required

## Slide 50: Bridging – step 3
Cover both the wound and foam bridge with Transparent Film
Film should extend 2in (5.1 cm) beyond wound margin to facilitate adequate seal

## Slide 51: Bridging – step 4
Cut a hole in the center of film over bridge foam
The hole should be at least .79in (2cm) in diameter
Remove any excess trimmed film
The center opening of the RENASYS™ Soft Port Dressing will be aligned over the hole to deliver negative pressure

## Slide 52: Bridging – step 5
Initiate therapy and ensure that the entire dressing including the RENASYS™ Soft Port is collapsed
Finished and leak free dressing should be firm to the touch
Carefully monitor patients for signs of bleeding, which may lead to interruption in therapy and hemodynamic instability. If such symptoms are observed, immediately discontinue therapy, take appropriate measures to control bleeding, and contact treating clinician

## Slide 53: Optimization of therapy: Bridging
Connecting two wounds of different etiologies via a bridge is not recommended due to the risk of cross contamination
Connecting two wounds of differing infection status via a bridge (non-infected with infected) is not recommended due to the risk of cross contamination
When using a bridging technique, choose a location where the bridge is least likely to be weight bearing. This will support the bridge remaining open to optimize transfer of fluid and vacuum
To optimize transfer of fluid and vacuum through a bridge, foam is recommended. In the instance that gauze is used, particularly for wounds with moderate to heavy drainage, more frequent inspection of the dressing and dressing changes may be required

## Slide 54: Changing canister
Light flashes yellow
LCD Warning! Blockage/Full
Canister is full
Alarm will sound - beeps 2 times every 20 seconds
Note: To optimize canister over-capacity alarm functionality, canister must be changed in the event the canister contents reach maximum volume indication (300mL or 750mL fill line).
Do not wait for alarm activation to change the canister
Notes: Complete blockage/canister over-capacity alarm:
The system detects a complete blockage in the system, this includes a canister where contents have reached
maximum volume capacity. There will be no negative pressure at the wound site. Once corrected, alarm will
Automatically reset and status/alarm indicator will return to solid green.
Cautions – Complete blockage/canister over-capacity alarm:
Complete blockage/canister over-capacity alarm activation is driven by occlusion of filter located at top of canister.
As canister fills with fluid, the filter will begin to occlude as fluid makes continuous contact, creating a complete
blockage in vacuum circuit. A completely occluded filter will result in canister over-capacity alarm activation.
Complete blockage/canister over-capacity alarm will not activate until filter is completely occluded. Canister
orientation, rate at which fluid enters canister and how exudate solidifies can impact filter occlusion and canister
over-capacity alarm functionality.
Canisters should be kept in an upright orientation to optimize canister over-capacity alarm functionality while
maximizing canister volume. When canister is operating in horizontal orientation, facing upward, alarm functionality
may be compromised.
To optimize canister over-capacity alarm functionality, canister must be changed in the event the canister contents
reach maximum volume indication (300mL or 750mL fill line).
Do not wait for alarm activation to change the canister.

## Slide 55: Changing canister
To change canister:
Clamp dressing tubing
Stop therapy
Disconnect canister tube from dressing tube and cap both tubing connectors
Open orange canister clips
Notes: Removing or changing canister
1. Hold orange Quick Click Connector above wound to
help ensure exudate does not leak from tubing.
2. Turn device off.
3. Disconnect canister tubing from Soft Port.
4. Insert tethered cap into both sides of orange Quick
Click Connector.
5. Release orange canister clips on sides of canister
and gently pull away from device.

## Slide 56: Changing canister
Removing canister:
Gently pull canister in downward motion to release from device

## Slide 57: Changing canister
Attaching new canister:
Ensure device is off
Open both orange clips on the sides of canister
Align canister so that viewing window is facing forward
Push canister gently over device inlet port
Engage both canister clips. Clips will click when properly engaged
Resume therapy

## Slide 58: Following canister installation, verify functional complete blockage alarm:
Turn device on and insert tethered cap of canister tubing into connector to simulate a blockage
A functional device will activate a complete blockage alarm within 5 minutes
In the case that complete blockage alarm does not activate, check canister installation and contact your Smith & Nephew authorized representative

## Slide 59: RENASYS™ GO Quick Reference Guide (hang tag)
Notes: Each RENASYS GO pump will have a hang tag as a quick reference for pump functionality and troubleshooting and other information

## Slide 60: Dressing removal
When disconnecting the Quick Click Connector, protect the tube ends by inserting the tethered caps immediately before turning the device off
As with all adhesive products, apply and remove the dressing carefully from sensitive skin to avoid skin stripping, especially after frequent dressing changes
Ensure all wound filler material placed in wound has been removed before dressing wound
If dressing adheres to wound, apply normal saline onto wound dressing and let set for 15-30 minutes before gently removing
Appropriately discard used wound dressings observing facility protocol for medical waste handling

## Slide 61: Know when to stop or change treatment
Established goal of therapy has been met
No improvement or reduction in wound volume has been documented consecutively for 2 weeks
Individual patient and wound considerations may vary clinical decision to discontinue NPWT

## Slide 62: NPWT Product and clinical support hotline 1-800-876-1261
24/7 Hotline for NPWT clinical support
Staffed by clinicians
Skilled in trouble shooting and product support
Provide support to clinicians, patients, and family members
First point of contact for dressing application
Trouble shooting
General product information

## Slide 63
©2016 Smith & Nephew, Inc. All rights reserved.
™Trademark of Smith & Nephew.
For detailed product information, including indications for use, contraindications, precautions and warnings, please consult the product’s Instructions for Use (IFU) prior to use.
Questions
RNPE-32-0716-UE